## Sheet
PERFORMANCE TEST CARD (AS RECEIVED AND AS COMPLETED)
| DATE: |  |  |  | REPORT NUMBER:
| MODEL NUMBER: | MSO44,MSO46 |  |  | DESCRIPTION: | 1 GHz Oscilloscope
| SERIAL NUMBER: |  |  |  | MANUFACTURER: | Tektronix
| CUSTOMER ID CODE: |  |  |  | OPTIONS: | Nil
| HUMIDITY: (%RH) | 30 - 60 |  |  | TESTED BY:
| TEMPERATURE: (Deg C) | 23 ± 2 |  |  | REFERENCE NO.: | RFTS_TEKMSO46
| ISSUE: | A.00.00
Channels
Self Test
| Test Conditions | Pass/Fail | Pass/Fail
| Self Test | PASS/FAIL | [6 empty] | PASS/FAIL
Signal Path Compensation
| Test Conditions | Pass/Fail | Pass/Fail
| Self Test | PASS/FAIL
[13 empty] | Test | Channel | Coupling | Scale | Voltage | Offset | Bandwidth | Impedance | Invert
Input Impedance
| Test Conditions | Vertical Scale | Lower Limit | Measured | Upper Limit | Units | Pass/Fail
| Ch1, 1 MOhm | 100 mV/Div | 0.99 |  | 1.01 | MOhm | [6 empty] | IMP | 1 | DC | 0.1 |  |  |  | 1M
| Ch1, 50Ohm | 10 mV/Div | 49.5 |  | 50.5 | Ohm | [6 empty] | IMP | 1 | DC | 0.01 |  |  |  | 50
| Ch1, 50Ohm | 100 mV/Div | 49.5 |  | 50.5 | Ohm | [6 empty] | IMP | 1 | DC | 0.1 |  |  |  | 50
| Ch2, 1 MOhm | 100 mV/Div | 0.99 |  | 1.01 | MOhm | [6 empty] | IMP | 2 | DC | 0.1 |  |  |  | 1M
| Ch2, 50Ohm | 10 mV/Div | 49.5 |  | 50.5 | Ohm | [6 empty] | IMP | 2 | DC | 0.01 |  |  |  | 50
| Ch2, 50Ohm | 100 mV/Div | 49.5 |  | 50.5 | Ohm | [6 empty] | IMP | 2 | DC | 0.1 |  |  |  | 50
| Ch3, 1 MOhm | 100 mV/Div | 0.99 |  | 1.01 | MOhm | [6 empty] | IMP | 3 | DC | 0.1 |  |  |  | 1M
| Ch3, 50Ohm | 10 mV/Div | 49.5 |  | 50.5 | Ohm | [6 empty] | IMP | 3 | DC | 0.01 |  |  |  | 50
| Ch3, 50Ohm | 100 mV/Div | 49.5 |  | 50.5 | Ohm | [6 empty] | IMP | 3 | DC | 0.1 |  |  |  | 50
| Ch4, 1 MOhm | 100 mV/Div | 0.99 |  | 1.01 | MOhm | [6 empty] | IMP | 4 | DC | 0.1 |  |  |  | 1M
| Ch4, 50Ohm | 10 mV/Div | 49.5 |  | 50.5 | Ohm | [6 empty] | IMP | 4 | DC | 0.01 |  |  |  | 50
| Ch4, 50Ohm | 100 mV/Div | 49.5 |  | 50.5 | Ohm | [6 empty] | IMP | 4 | DC | 0.1 |  |  |  | 50
6 | Ch5, 1 MOhm | 100 mV/Div | 0.99 |  | 1.01 | MOhm | [6 empty] | IMP | 5 | DC | 0.1 |  |  |  | 1M
6 | Ch5, 50Ohm | 10 mV/Div | 49.5 |  | 50.5 | Ohm | [6 empty] | IMP | 5 | DC | 0.01 |  |  |  | 50
6 | Ch5, 50Ohm | 100 mV/Div | 49.5 |  | 50.5 | Ohm | [6 empty] | IMP | 5 | DC | 0.1 |  |  |  | 50
6 | Ch6, 1 MOhm | 100 mV/Div | 0.99 |  | 1.01 | MOhm | [6 empty] | IMP | 6 | DC | 0.1 |  |  |  | 1M
6 | Ch6, 50Ohm | 10 mV/Div | 49.5 |  | 50.5 | Ohm | [6 empty] | IMP | 6 | DC | 0.01 |  |  |  | 50
6 | Ch6, 50Ohm | 100 mV/Div | 49.5 |  | 50.5 | Ohm | [6 empty] | IMP | 6 | DC | 0.1 |  |  |  | 50
DC Gain Accuracy
| Test Conditions | Vertical Scale | Lower Limit | Measured | Upper Limit | Units | Pass/Fail |  | Voltage (units) | Limit | Percent | Multiplier | Test | Channel | Coupling | Scale | Voltage | Offset | Bandwidth | Impedance | Invert
| Ch 1 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 1 mV/div |  |  |  | mV |  |  |  |  | 2 |  | DCV-BAL | 1 | DC | 0.001 | 0.0045 | 0 | 20M | 1M
| Ch 1 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 2 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 1 | DC | 0.002 | 0.009 | 0 | 20M | 1M
| Ch 1 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 5 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 1 | DC | 0.005 | 0.0225 | 0 | 20M | 1M
| Ch 1 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 10 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 1 | DC | 0.01 | 0.045 | 0 | 20M | 1M
| Ch 1 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 20 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 1 | DC | 0.02 | 0.09 | 0 | 20M | 1M
| Ch 1 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 50 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 1 | DC | 0.05 | 0.225 | 0 | 20M | 1M
| Ch 1 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 100 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 1 | DC | 0.1 | 0.45 | 0 | 20M | 1M
| Ch 1 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 200 mV/div |  |  |  | V |  |  |  |  | 1 |  | DCV-BAL | 1 | DC | 0.2 | 0.9 | 0 | 20M | 1M
| Ch 1 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 500 mV/div |  |  |  | V |  |  |  |  | 1 |  | DCV-BAL | 1 | DC | 0.5 | 2.25 | 0 | 20M | 1M
| Ch 1 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 1 V/div |  |  |  | V |  |  |  |  | 1 |  | DCV-BAL | 1 | DC | 1 | 4.5 | 0 | 20M | 1M
| Ch 1 0V Offset, 1M Ohm, 0V Vert, 250 MHz BW | 20 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 1 | DC | 0.02 | 0.09 | 0 | 250M | 1M
| Ch 1 0V Offset, 1M Ohm, 0V Vert, Full BW | 20 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 1 | DC | 0.02 | 0.09 | 0 | FULL | 1M
| Ch 1 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 1 mV/div |  |  |  | mV |  |  |  |  | 2.5 |  | DCV-BAL | 1 | DC | 0.001 | 0.0045 | 0 | 20M | 50
| Ch 1 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 2 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 1 | DC | 0.002 | 0.009 | 0 | 20M | 50
| Ch 1 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 5 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 1 | DC | 0.005 | 0.0225 | 0 | 20M | 50
| Ch 1 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 10 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 1 | DC | 0.01 | 0.045 | 0 | 20M | 50
| Ch 1 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 20 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 1 | DC | 0.02 | 0.09 | 0 | 20M | 50
| Ch 1 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 50 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 1 | DC | 0.05 | 0.225 | 0 | 20M | 50
| Ch 1 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 100 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 1 | DC | 0.1 | 0.45 | 0 | 20M | 50
| Ch 1 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 200 mV/div |  |  |  | V |  |  |  |  | 1 |  | DCV-BAL | 1 | DC | 0.2 | 0.9 | 0 | 20M | 50
| Ch 1 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 500 mV/div |  |  |  | V |  |  |  |  | 1 |  | DCV-BAL | 1 | DC | 0.5 | 2.25 | 0 | 20M | 50
| Ch 1 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 1 V/div |  |  |  | V |  |  |  |  | 1 |  | DCV-BAL | 1 | DC | 1 | 3 | 0 | 20M | 50
| Ch 1 0V Offset, 50 Ohm, 0V Vert, 250 MHz BW | 20 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 1 | DC | 0.02 | 0.09 | 0 | 250M | 50
| Ch 1 0V Offset, 50 Ohm, 0V Vert, Full BW | 20 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 1 | DC | 0.02 | 0.09 | 0 | FULL | 50
| Ch 2 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 1 mV/div |  |  |  | mV |  |  |  |  | 2 |  | DCV-BAL | 2 | DC | 0.001 | 0.0045 | 0 | 20M | 1M
| Ch 2 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 2 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 2 | DC | 0.002 | 0.009 | 0 | 20M | 1M
| Ch 2 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 5 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 2 | DC | 0.005 | 0.0225 | 0 | 20M | 1M
| Ch 2 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 10 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 2 | DC | 0.01 | 0.045 | 0 | 20M | 1M
| Ch 2 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 20 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 2 | DC | 0.02 | 0.09 | 0 | 20M | 1M
| Ch 2 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 50 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 2 | DC | 0.05 | 0.225 | 0 | 20M | 1M
| Ch 2 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 100 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 2 | DC | 0.1 | 0.45 | 0 | 20M | 1M
| Ch 2 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 200 mV/div |  |  |  | V |  |  |  |  | 1 |  | DCV-BAL | 2 | DC | 0.2 | 0.9 | 0 | 20M | 1M
| Ch 2 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 500 mV/div |  |  |  | V |  |  |  |  | 1 |  | DCV-BAL | 2 | DC | 0.5 | 2.25 | 0 | 20M | 1M
| Ch 2 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 1 V/div |  |  |  | V |  |  |  |  | 1 |  | DCV-BAL | 2 | DC | 1 | 4.5 | 0 | 20M | 1M
| Ch 2 0V Offset, 1M Ohm, 0V Vert, 250 MHz BW | 20 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 2 | DC | 0.02 | 0.09 | 0 | 250M | 1M
| Ch 2 0V Offset, 1M Ohm, 0V Vert, Full BW | 20 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 2 | DC | 0.02 | 0.09 | 0 | FULL | 1M
| Ch 2 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 1 mV/div |  |  |  | mV |  |  |  |  | 2.5 |  | DCV-BAL | 2 | DC | 0.001 | 0.0045 | 0 | 20M | 50
| Ch 2 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 2 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 2 | DC | 0.002 | 0.009 | 0 | 20M | 50
| Ch 2 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 5 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 2 | DC | 0.005 | 0.0225 | 0 | 20M | 50
| Ch 2 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 10 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 2 | DC | 0.01 | 0.045 | 0 | 20M | 50
| Ch 2 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 20 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 2 | DC | 0.02 | 0.09 | 0 | 20M | 50
| Ch 2 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 50 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 2 | DC | 0.05 | 0.225 | 0 | 20M | 50
| Ch 2 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 100 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 2 | DC | 0.1 | 0.45 | 0 | 20M | 50
| Ch 2 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 200 mV/div |  |  |  | V |  |  |  |  | 1 |  | DCV-BAL | 2 | DC | 0.2 | 0.9 | 0 | 20M | 50
| Ch 2 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 500 mV/div |  |  |  | V |  |  |  |  | 1 |  | DCV-BAL | 2 | DC | 0.5 | 2.25 | 0 | 20M | 50
| Ch 2 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 1 V/div |  |  |  | V |  |  |  |  | 1 |  | DCV-BAL | 2 | DC | 1 | 3 | 0 | 20M | 50
| Ch 2 0V Offset, 50 Ohm, 0V Vert, 250 MHz BW | 20 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 2 | DC | 0.02 | 0.09 | 0 | 250M | 50
| Ch 2 0V Offset, 50 Ohm, 0V Vert, Full BW | 20 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 2 | DC | 0.02 | 0.09 | 0 | FULL | 50
| Ch 3 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 1 mV/div |  |  |  | mV |  |  |  |  | 2 |  | DCV-BAL | 3 | DC | 0.001 | 0.0045 | 0 | 20M | 1M
| Ch 3 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 2 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 3 | DC | 0.002 | 0.009 | 0 | 20M | 1M
| Ch 3 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 5 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 3 | DC | 0.005 | 0.0225 | 0 | 20M | 1M
| Ch 3 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 10 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 3 | DC | 0.01 | 0.045 | 0 | 20M | 1M
| Ch 3 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 20 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 3 | DC | 0.02 | 0.09 | 0 | 20M | 1M
| Ch 3 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 50 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 3 | DC | 0.05 | 0.225 | 0 | 20M | 1M
| Ch 3 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 100 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 3 | DC | 0.1 | 0.45 | 0 | 20M | 1M
| Ch 3 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 200 mV/div |  |  |  | V |  |  |  |  | 1 |  | DCV-BAL | 3 | DC | 0.2 | 0.9 | 0 | 20M | 1M
| Ch 3 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 500 mV/div |  |  |  | V |  |  |  |  | 1 |  | DCV-BAL | 3 | DC | 0.5 | 2.25 | 0 | 20M | 1M
| Ch 3 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 1 V/div |  |  |  | V |  |  |  |  | 1 |  | DCV-BAL | 3 | DC | 1 | 4.5 | 0 | 20M | 1M
| Ch 3 0V Offset, 1M Ohm, 0V Vert, 250 MHz BW | 20 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 3 | DC | 0.02 | 0.09 | 0 | 250M | 1M
| Ch 3 0V Offset, 1M Ohm, 0V Vert, Full BW | 20 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 3 | DC | 0.02 | 0.09 | 0 | FULL | 1M
| Ch 3 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 1 mV/div |  |  |  | mV |  |  |  |  | 2.5 |  | DCV-BAL | 3 | DC | 0.001 | 0.0045 | 0 | 20M | 50
| Ch 3 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 2 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 3 | DC | 0.002 | 0.009 | 0 | 20M | 50
| Ch 3 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 5 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 3 | DC | 0.005 | 0.0225 | 0 | 20M | 50
| Ch 3 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 10 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 3 | DC | 0.01 | 0.045 | 0 | 20M | 50
| Ch 3 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 20 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 3 | DC | 0.02 | 0.09 | 0 | 20M | 50
| Ch 3 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 50 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 3 | DC | 0.05 | 0.225 | 0 | 20M | 50
| Ch 3 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 100 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 3 | DC | 0.1 | 0.45 | 0 | 20M | 50
| Ch 3 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 200 mV/div |  |  |  | V |  |  |  |  | 1 |  | DCV-BAL | 3 | DC | 0.2 | 0.9 | 0 | 20M | 50
| Ch 3 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 500 mV/div |  |  |  | V |  |  |  |  | 1 |  | DCV-BAL | 3 | DC | 0.5 | 2.25 | 0 | 20M | 50
| Ch 3 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 1 V/div |  |  |  | V |  |  |  |  | 1 |  | DCV-BAL | 3 | DC | 1 | 3 | 0 | 20M | 50
| Ch 3 0V Offset, 50 Ohm, 0V Vert, 250 MHz BW | 20 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 3 | DC | 0.02 | 0.09 | 0 | 250M | 50
| Ch 3 0V Offset, 50 Ohm, 0V Vert, Full BW | 20 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 3 | DC | 0.02 | 0.09 | 0 | FULL | 50
| Ch 4 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 1 mV/div |  |  |  | mV |  |  |  |  | 2 |  | DCV-BAL | 4 | DC | 0.001 | 0.0045 | 0 | 20M | 1M
| Ch 4 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 2 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 4 | DC | 0.002 | 0.009 | 0 | 20M | 1M
| Ch 4 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 5 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 4 | DC | 0.005 | 0.0225 | 0 | 20M | 1M
| Ch 4 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 10 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 4 | DC | 0.01 | 0.045 | 0 | 20M | 1M
| Ch 4 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 20 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 4 | DC | 0.02 | 0.09 | 0 | 20M | 1M
| Ch 4 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 50 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 4 | DC | 0.05 | 0.225 | 0 | 20M | 1M
| Ch 4 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 100 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 4 | DC | 0.1 | 0.45 | 0 | 20M | 1M
| Ch 4 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 200 mV/div |  |  |  | V |  |  |  |  | 1 |  | DCV-BAL | 4 | DC | 0.2 | 0.9 | 0 | 20M | 1M
| Ch 4 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 500 mV/div |  |  |  | V |  |  |  |  | 1 |  | DCV-BAL | 4 | DC | 0.5 | 2.25 | 0 | 20M | 1M
| Ch 4 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 1 V/div |  |  |  | V |  |  |  |  | 1 |  | DCV-BAL | 4 | DC | 1 | 4.5 | 0 | 20M | 1M
| Ch 4 0V Offset, 1M Ohm, 0V Vert, 250 MHz BW | 20 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 4 | DC | 0.02 | 0.09 | 0 | 250M | 1M
| Ch 4 0V Offset, 1M Ohm, 0V Vert, Full BW | 20 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 4 | DC | 0.02 | 0.09 | 0 | FULL | 1M
| Ch 4 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 1 mV/div |  |  |  | mV |  |  |  |  | 2.5 |  | DCV-BAL | 4 | DC | 0.001 | 0.0045 | 0 | 20M | 50
| Ch 4 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 2 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 4 | DC | 0.002 | 0.009 | 0 | 20M | 50
| Ch 4 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 5 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 4 | DC | 0.005 | 0.0225 | 0 | 20M | 50
| Ch 4 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 10 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 4 | DC | 0.01 | 0.045 | 0 | 20M | 50
| Ch 4 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 20 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 4 | DC | 0.02 | 0.09 | 0 | 20M | 50
| Ch 4 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 50 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 4 | DC | 0.05 | 0.225 | 0 | 20M | 50
| Ch 4 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 100 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 4 | DC | 0.1 | 0.45 | 0 | 20M | 50
| Ch 4 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 200 mV/div |  |  |  | V |  |  |  |  | 1 |  | DCV-BAL | 4 | DC | 0.2 | 0.9 | 0 | 20M | 50
| Ch 4 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 500 mV/div |  |  |  | V |  |  |  |  | 1 |  | DCV-BAL | 4 | DC | 0.5 | 2.25 | 0 | 20M | 50
| Ch 4 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 1 V/div |  |  |  | V |  |  |  |  | 1 |  | DCV-BAL | 4 | DC | 1 | 3 | 0 | 20M | 50
| Ch 4 0V Offset, 50 Ohm, 0V Vert, 250 MHz BW | 20 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 4 | DC | 0.02 | 0.09 | 0 | 250M | 50
| Ch 4 0V Offset, 50 Ohm, 0V Vert, Full BW | 20 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 4 | DC | 0.02 | 0.09 | 0 | FULL | 50
6 | Ch 5 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 1 mV/div |  |  |  | mV |  |  |  |  | 2 |  | DCV-BAL | 5 | DC | 0.001 | 0.0045 | 0 | 20M | 1M
6 | Ch 5 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 2 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 5 | DC | 0.002 | 0.009 | 0 | 20M | 1M
6 | Ch 5 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 5 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 5 | DC | 0.005 | 0.0225 | 0 | 20M | 1M
6 | Ch 5 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 10 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 5 | DC | 0.01 | 0.045 | 0 | 20M | 1M
6 | Ch 5 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 20 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 5 | DC | 0.02 | 0.09 | 0 | 20M | 1M
6 | Ch 5 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 50 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 5 | DC | 0.05 | 0.225 | 0 | 20M | 1M
6 | Ch 5 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 100 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 5 | DC | 0.1 | 0.45 | 0 | 20M | 1M
6 | Ch 5 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 200 mV/div |  |  |  | V |  |  |  |  | 1 |  | DCV-BAL | 5 | DC | 0.2 | 0.9 | 0 | 20M | 1M
6 | Ch 5 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 500 mV/div |  |  |  | V |  |  |  |  | 1 |  | DCV-BAL | 5 | DC | 0.5 | 2.25 | 0 | 20M | 1M
6 | Ch 5 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 1 V/div |  |  |  | V |  |  |  |  | 1 |  | DCV-BAL | 5 | DC | 1 | 4.5 | 0 | 20M | 1M
6 | Ch 5 0V Offset, 1M Ohm, 0V Vert, 250 MHz BW | 20 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 5 | DC | 0.02 | 0.09 | 0 | 250M | 1M
6 | Ch 5 0V Offset, 1M Ohm, 0V Vert, Full BW | 20 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 5 | DC | 0.02 | 0.09 | 0 | FULL | 1M
6 | Ch 5 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 1 mV/div |  |  |  | mV |  |  |  |  | 2.5 |  | DCV-BAL | 5 | DC | 0.001 | 0.0045 | 0 | 20M | 50
6 | Ch 5 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 2 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 5 | DC | 0.002 | 0.009 | 0 | 20M | 50
6 | Ch 5 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 5 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 5 | DC | 0.005 | 0.0225 | 0 | 20M | 50
6 | Ch 5 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 10 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 5 | DC | 0.01 | 0.045 | 0 | 20M | 50
6 | Ch 5 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 20 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 5 | DC | 0.02 | 0.09 | 0 | 20M | 50
6 | Ch 5 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 50 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 5 | DC | 0.05 | 0.225 | 0 | 20M | 50
6 | Ch 5 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 100 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 5 | DC | 0.1 | 0.45 | 0 | 20M | 50
6 | Ch 5 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 200 mV/div |  |  |  | V |  |  |  |  | 1 |  | DCV-BAL | 5 | DC | 0.2 | 0.9 | 0 | 20M | 50
6 | Ch 5 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 500 mV/div |  |  |  | V |  |  |  |  | 1 |  | DCV-BAL | 5 | DC | 0.5 | 2.25 | 0 | 20M | 50
6 | Ch 5 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 1 V/div |  |  |  | V |  |  |  |  | 1 |  | DCV-BAL | 5 | DC | 1 | 3 | 0 | 20M | 50
6 | Ch 5 0V Offset, 50 Ohm, 0V Vert, 250 MHz BW | 20 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 5 | DC | 0.02 | 0.09 | 0 | 250M | 50
6 | Ch 5 0V Offset, 50 Ohm, 0V Vert, Full BW | 20 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 5 | DC | 0.02 | 0.09 | 0 | FULL | 50
6 | Ch 6 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 1 mV/div |  |  |  | mV |  |  |  |  | 2 |  | DCV-BAL | 6 | DC | 0.001 | 0.0045 | 0 | 20M | 1M
6 | Ch 6 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 2 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 6 | DC | 0.002 | 0.009 | 0 | 20M | 1M
6 | Ch 6 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 5 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 6 | DC | 0.005 | 0.0225 | 0 | 20M | 1M
6 | Ch 6 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 10 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 6 | DC | 0.01 | 0.045 | 0 | 20M | 1M
6 | Ch 6 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 20 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 6 | DC | 0.02 | 0.09 | 0 | 20M | 1M
6 | Ch 6 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 50 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 6 | DC | 0.05 | 0.225 | 0 | 20M | 1M
6 | Ch 6 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 100 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 6 | DC | 0.1 | 0.45 | 0 | 20M | 1M
6 | Ch 6 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 200 mV/div |  |  |  | V |  |  |  |  | 1 |  | DCV-BAL | 6 | DC | 0.2 | 0.9 | 0 | 20M | 1M
6 | Ch 6 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 500 mV/div |  |  |  | V |  |  |  |  | 1 |  | DCV-BAL | 6 | DC | 0.5 | 2.25 | 0 | 20M | 1M
6 | Ch 6 0V Offset, 1M Ohm, 0V Vert, 20 MHz BW | 1 V/div |  |  |  | V |  |  |  |  | 1 |  | DCV-BAL | 6 | DC | 1 | 4.5 | 0 | 20M | 1M
6 | Ch 6 0V Offset, 1M Ohm, 0V Vert, 250 MHz BW | 20 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 6 | DC | 0.02 | 0.09 | 0 | 250M | 1M
6 | Ch 6 0V Offset, 1M Ohm, 0V Vert, Full BW | 20 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 6 | DC | 0.02 | 0.09 | 0 | FULL | 1M
6 | Ch 6 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 1 mV/div |  |  |  | mV |  |  |  |  | 2.5 |  | DCV-BAL | 6 | DC | 0.001 | 0.0045 | 0 | 20M | 50
6 | Ch 6 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 2 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 6 | DC | 0.002 | 0.009 | 0 | 20M | 50
6 | Ch 6 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 5 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 6 | DC | 0.005 | 0.0225 | 0 | 20M | 50
6 | Ch 6 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 10 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 6 | DC | 0.01 | 0.045 | 0 | 20M | 50
6 | Ch 6 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 20 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 6 | DC | 0.02 | 0.09 | 0 | 20M | 50
6 | Ch 6 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 50 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 6 | DC | 0.05 | 0.225 | 0 | 20M | 50
6 | Ch 6 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 100 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 6 | DC | 0.1 | 0.45 | 0 | 20M | 50
6 | Ch 6 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 200 mV/div |  |  |  | V |  |  |  |  | 1 |  | DCV-BAL | 6 | DC | 0.2 | 0.9 | 0 | 20M | 50
6 | Ch 6 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 500 mV/div |  |  |  | V |  |  |  |  | 1 |  | DCV-BAL | 6 | DC | 0.5 | 2.25 | 0 | 20M | 50
6 | Ch 6 0V Offset, 50 Ohm, 0V Vert, 20 MHz BW | 1 V/div |  |  |  | V |  |  |  |  | 1 |  | DCV-BAL | 6 | DC | 1 | 3 | 0 | 20M | 50
6 | Ch 6 0V Offset, 50 Ohm, 0V Vert, 250 MHz BW | 20 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 6 | DC | 0.02 | 0.09 | 0 | 250M | 50
6 | Ch 6 0V Offset, 50 Ohm, 0V Vert, Full BW | 20 mV/div |  |  |  | mV |  |  |  |  | 1 |  | DCV-BAL | 6 | DC | 0.02 | 0.09 | 0 | FULL | 50
DC Offset
| Test Conditions | Vertical Scale | Vertical Offset | Lower Limit | Measured | Upper Limit | Units | Pass/Fail
| Ch 1, 1M Ohm, 20 MHz BW | 1 mV/div | 900 mV | 895.3 |  | 904.7 | mV | [5 empty] | DCV | 1 | DC | 0.001 | 0.9 | 0.9 | 20M | 1M
| Ch 1, 1M Ohm, 20 MHz BW | 1 mV/div | 0 mV | -0.2 |  | 0.2 | mV | [5 empty] | DCV | 1 | DC | 0.001 | 0 | 0 | 20M | 1M
| Ch 1, 1M Ohm, 20 MHz BW | 1 mV/div | -900 mV | -904.7 |  | -895.3 | mV | [5 empty] | DCV | 1 | DC | 0.001 | -0.9 | -0.9 | 20M | 1M
| Ch 1, 1M Ohm, 20 MHz BW | 100 mV/div | 1.0 V | 0.975 |  | 1.025 | V | [5 empty] | DCV | 1 | DC | 0.1 | 1 | 1 | 20M | 1M
| Ch 1, 1M Ohm, 20 MHz BW | 100 mV/div | 0 V | -0.02 |  | 0.02 | V | [5 empty] | DCV | 1 | DC | 0.1 | 0 | 0 | 20M | 1M
| Ch 1, 1M Ohm, 20 MHz BW | 100 mV/div | -1.0 V | -1.025 |  | -0.975 | V | [5 empty] | DCV | 1 | DC | 0.1 | -1 | -1 | 20M | 1M
| Ch 1, 1M Ohm, 20 MHz BW | 500 mV/div | 9.0 V | 8.855 |  | 9.145 | V | [5 empty] | DCV | 1 | DC | 0.5 | 9 | 9 | 20M | 1M
| Ch 1, 1M Ohm, 20 MHz BW | 500 mV/div | 0 V | -0.1 |  | 0.1 | V | [5 empty] | DCV | 1 | DC | 0.5 | 0 | 0 | 20M | 1M
| Ch 1, 1M Ohm, 20 MHz BW | 500 mV/div | -9.0 V | -9.145 |  | -8.855 | V | [5 empty] | DCV | 1 | DC | 0.5 | -9 | -9 | 20M | 1M
| Ch 1, 1M Ohm, 20 MHz BW | 5 V/div | 10 | 8.95 |  | 11.05 | V | [5 empty] | DCV | 1 | DC | 5 | 10 | 10 | 20M | 1M
| Ch 1, 1M Ohm, 20 MHz BW | 5 V/div | 0 V | -1 |  | 1 | V | [5 empty] | DCV | 1 | DC | 5 | 0 | 0 | 20M | 1M
| Ch 1, 1M Ohm, 20 MHz BW | 5 V/div | -10 V | -11.05 |  | -8.95 | V | [5 empty] | DCV | 1 | DC | 5 | -10 | -10 | 20M | 1M
| Ch 1, 50 Ohm, 20 MHz BW | 1 mV/div | 900 mV | 895.3 |  | 904.7 | mV | [5 empty] | DCV | 1 | DC | 0.001 | 0.9 | 0.9 | 20M | 50
| Ch 1, 50 Ohm, 20 MHz BW | 1 mV/div | 0 mV | -0.2 |  | 0.2 | mV | [5 empty] | DCV | 1 | DC | 0.001 | 0 | 0 | 20M | 50
| Ch 1, 50 Ohm, 20 MHz BW | 1 mV/div | -900 mV | -904.7 |  | -895.3 | mV | [5 empty] | DCV | 1 | DC | 0.001 | -0.9 | -0.9 | 20M | 50
| Ch 1, 50 Ohm, 20 MHz BW | 100 mV/div | 3 V | 2.97 |  | 3.03 | V | [5 empty] | DCV | 1 | DC | 0.1 | 3 | 3 | 20M | 50
| Ch 1, 50 Ohm, 20 MHz BW | 100 mV/div | 0 V | -0.02 |  | 0.02 | V | [5 empty] | DCV | 1 | DC | 0.1 | 0 | 0 | 20M | 50
| Ch 1, 50 Ohm, 20 MHz BW | 100 mV/div | -3 V | -3.03 |  | -2.97 | V | [5 empty] | DCV | 1 | DC | 0.1 | -3 | -3 | 20M | 50
| Ch 2, 1M Ohm, 20 MHz BW | 1 mV/div | 900 mV | 895.3 |  | 904.7 | mV | [5 empty] | DCV | 2 | DC | 0.001 | 0.9 | 0.9 | 20M | 1M
| Ch 2, 1M Ohm, 20 MHz BW | 1 mV/div | 0 mV | -0.2 |  | 0.2 | mV | [5 empty] | DCV | 2 | DC | 0.001 | 0 | 0 | 20M | 1M
| Ch 2, 1M Ohm, 20 MHz BW | 1 mV/div | -900 mV | -904.7 |  | -895.3 | mV | [5 empty] | DCV | 2 | DC | 0.001 | -0.9 | -0.9 | 20M | 1M
| Ch 2, 1M Ohm, 20 MHz BW | 100 mV/div | 1.0 V | 0.975 |  | 1.025 | V | [5 empty] | DCV | 2 | DC | 0.1 | 1 | 1 | 20M | 1M
| Ch 2, 1M Ohm, 20 MHz BW | 100 mV/div | 0 V | -0.02 |  | 0.02 | V | [5 empty] | DCV | 2 | DC | 0.1 | 0 | 0 | 20M | 1M
| Ch 2, 1M Ohm, 20 MHz BW | 100 mV/div | -1.0 V | -1.025 |  | -0.975 | V | [5 empty] | DCV | 2 | DC | 0.1 | -1 | -1 | 20M | 1M
| Ch 2, 1M Ohm, 20 MHz BW | 500 mV/div | 9.0 V | 8.855 |  | 9.145 | V | [5 empty] | DCV | 2 | DC | 0.5 | 9 | 9 | 20M | 1M
| Ch 2, 1M Ohm, 20 MHz BW | 500 mV/div | 0 V | -0.1 |  | 0.1 | V | [5 empty] | DCV | 2 | DC | 0.5 | 0 | 0 | 20M | 1M
| Ch 2, 1M Ohm, 20 MHz BW | 500 mV/div | -9.0 V | -9.145 |  | -8.855 | V | [5 empty] | DCV | 2 | DC | 0.5 | -9 | -9 | 20M | 1M
| Ch 2, 1M Ohm, 20 MHz BW | 5 V/div | 10 | 8.95 |  | 11.05 | V | [5 empty] | DCV | 2 | DC | 5 | 10 | 10 | 20M | 1M
| Ch 2, 1M Ohm, 20 MHz BW | 5 V/div | 0 V | -1 |  | 1 | V | [5 empty] | DCV | 2 | DC | 5 | 0 | 0 | 20M | 1M
| Ch 2, 1M Ohm, 20 MHz BW | 5 V/div | -10 V | -11.05 |  | -8.95 | V | [5 empty] | DCV | 2 | DC | 5 | -10 | -10 | 20M | 1M
| Ch 2, 50 Ohm, 20 MHz BW | 1 mV/div | 900 mV | 895.3 |  | 904.7 | mV | [5 empty] | DCV | 2 | DC | 0.001 | 0.9 | 0.9 | 20M | 50
| Ch 2, 50 Ohm, 20 MHz BW | 1 mV/div | 0 mV | -0.2 |  | 0.2 | mV | [5 empty] | DCV | 2 | DC | 0.001 | 0 | 0 | 20M | 50
| Ch 2, 50 Ohm, 20 MHz BW | 1 mV/div | -900 mV | -904.7 |  | -895.3 | mV | [5 empty] | DCV | 2 | DC | 0.001 | -0.9 | -0.9 | 20M | 50
| Ch 2, 50 Ohm, 20 MHz BW | 100 mV/div | 3 V | 2.97 |  | 3.03 | V | [5 empty] | DCV | 2 | DC | 0.1 | 3 | 3 | 20M | 50
| Ch 2, 50 Ohm, 20 MHz BW | 100 mV/div | 0 V | -0.02 |  | 0.02 | V | [5 empty] | DCV | 2 | DC | 0.1 | 0 | 0 | 20M | 50
| Ch 2, 50 Ohm, 20 MHz BW | 100 mV/div | -3 V | -3.03 |  | -2.97 | V | [5 empty] | DCV | 2 | DC | 0.1 | -3 | -3 | 20M | 50
| Ch 3, 1M Ohm, 20 MHz BW | 1 mV/div | 900 mV | 895.3 |  | 904.7 | mV | [5 empty] | DCV | 3 | DC | 0.001 | 0.9 | 0.9 | 20M | 1M
| Ch 3, 1M Ohm, 20 MHz BW | 1 mV/div | 0 mV | -0.2 |  | 0.2 | mV | [5 empty] | DCV | 3 | DC | 0.001 | 0 | 0 | 20M | 1M
| Ch 3, 1M Ohm, 20 MHz BW | 1 mV/div | -900 mV | -904.7 |  | -895.3 | mV | [5 empty] | DCV | 3 | DC | 0.001 | -0.9 | -0.9 | 20M | 1M
| Ch 3, 1M Ohm, 20 MHz BW | 100 mV/div | 1.0 V | 0.975 |  | 1.025 | V | [5 empty] | DCV | 3 | DC | 0.1 | 1 | 1 | 20M | 1M
| Ch 3, 1M Ohm, 20 MHz BW | 100 mV/div | 0 V | -0.02 |  | 0.02 | V | [5 empty] | DCV | 3 | DC | 0.1 | 0 | 0 | 20M | 1M
| Ch 3, 1M Ohm, 20 MHz BW | 100 mV/div | -1.0 V | -1.025 |  | -0.975 | V | [5 empty] | DCV | 3 | DC | 0.1 | -1 | -1 | 20M | 1M
| Ch 3, 1M Ohm, 20 MHz BW | 500 mV/div | 9.0 V | 8.855 |  | 9.145 | V | [5 empty] | DCV | 3 | DC | 0.5 | 9 | 9 | 20M | 1M
| Ch 3, 1M Ohm, 20 MHz BW | 500 mV/div | 0 V | -0.1 |  | 0.1 | V | [5 empty] | DCV | 3 | DC | 0.5 | 0 | 0 | 20M | 1M
| Ch 3, 1M Ohm, 20 MHz BW | 500 mV/div | -9.0 V | -9.145 |  | -8.855 | V | [5 empty] | DCV | 3 | DC | 0.5 | -9 | -9 | 20M | 1M
| Ch 3, 1M Ohm, 20 MHz BW | 5 V/div | 10 | 8.95 |  | 11.05 | V | [5 empty] | DCV | 3 | DC | 5 | 10 | 10 | 20M | 1M
| Ch 3, 1M Ohm, 20 MHz BW | 5 V/div | 0 V | -1 |  | 1 | V | [5 empty] | DCV | 3 | DC | 5 | 0 | 0 | 20M | 1M
| Ch 3, 1M Ohm, 20 MHz BW | 5 V/div | -10 V | -11.05 |  | -8.95 | V | [5 empty] | DCV | 3 | DC | 5 | -10 | -10 | 20M | 1M
| Ch 3, 50 Ohm, 20 MHz BW | 1 mV/div | 900 mV | 895.3 |  | 904.7 | mV | [5 empty] | DCV | 3 | DC | 0.001 | 0.9 | 0.9 | 20M | 50
| Ch 3, 50 Ohm, 20 MHz BW | 1 mV/div | 0 mV | -0.2 |  | 0.2 | mV | [5 empty] | DCV | 3 | DC | 0.001 | 0 | 0 | 20M | 50
| Ch 3, 50 Ohm, 20 MHz BW | 1 mV/div | -900 mV | -904.7 |  | -895.3 | mV | [5 empty] | DCV | 3 | DC | 0.001 | -0.9 | -0.9 | 20M | 50
| Ch 3, 50 Ohm, 20 MHz BW | 100 mV/div | 3 V | 2.97 |  | 3.03 | V | [5 empty] | DCV | 3 | DC | 0.1 | 3 | 3 | 20M | 50
| Ch 3, 50 Ohm, 20 MHz BW | 100 mV/div | 0 V | -0.02 |  | 0.02 | V | [5 empty] | DCV | 3 | DC | 0.1 | 0 | 0 | 20M | 50
| Ch 3, 50 Ohm, 20 MHz BW | 100 mV/div | -3 V | -3.03 |  | -2.97 | V | [5 empty] | DCV | 3 | DC | 0.1 | -3 | -3 | 20M | 50
| Ch 4, 1M Ohm, 20 MHz BW | 1 mV/div | 900 mV | 895.3 |  | 904.7 | mV | [5 empty] | DCV | 4 | DC | 0.001 | 0.9 | 0.9 | 20M | 1M
| Ch 4, 1M Ohm, 20 MHz BW | 1 mV/div | 0 mV | -0.2 |  | 0.2 | mV | [5 empty] | DCV | 4 | DC | 0.001 | 0 | 0 | 20M | 1M
| Ch 4, 1M Ohm, 20 MHz BW | 1 mV/div | -900 mV | -904.7 |  | -895.3 | mV | [5 empty] | DCV | 4 | DC | 0.001 | -0.9 | -0.9 | 20M | 1M
| Ch 4, 1M Ohm, 20 MHz BW | 100 mV/div | 1.0 V | 0.975 |  | 1.025 | V | [5 empty] | DCV | 4 | DC | 0.1 | 1 | 1 | 20M | 1M
| Ch 4, 1M Ohm, 20 MHz BW | 100 mV/div | 0 V | -0.02 |  | 0.02 | V | [5 empty] | DCV | 4 | DC | 0.1 | 0 | 0 | 20M | 1M
| Ch 4, 1M Ohm, 20 MHz BW | 100 mV/div | -1.0 V | -1.025 |  | -0.975 | V | [5 empty] | DCV | 4 | DC | 0.1 | -1 | -1 | 20M | 1M
| Ch 4, 1M Ohm, 20 MHz BW | 500 mV/div | 9.0 V | 8.855 |  | 9.145 | V | [5 empty] | DCV | 4 | DC | 0.5 | 9 | 9 | 20M | 1M
| Ch 4, 1M Ohm, 20 MHz BW | 500 mV/div | 0 V | -0.1 |  | 0.1 | V | [5 empty] | DCV | 4 | DC | 0.5 | 0 | 0 | 20M | 1M
| Ch 4, 1M Ohm, 20 MHz BW | 500 mV/div | -9.0 V | -9.145 |  | -8.855 | V | [5 empty] | DCV | 4 | DC | 0.5 | -9 | -9 | 20M | 1M
| Ch 4, 1M Ohm, 20 MHz BW | 5 V/div | 10 | 8.95 |  | 11.05 | V | [5 empty] | DCV | 4 | DC | 5 | 10 | 10 | 20M | 1M
| Ch 4, 1M Ohm, 20 MHz BW | 5 V/div | 0 V | -1 |  | 1 | V | [5 empty] | DCV | 4 | DC | 5 | 0 | 0 | 20M | 1M
| Ch 4, 1M Ohm, 20 MHz BW | 5 V/div | -10 V | -11.05 |  | -8.95 | V | [5 empty] | DCV | 4 | DC | 5 | -10 | -10 | 20M | 1M
| Ch 4, 50 Ohm, 20 MHz BW | 1 mV/div | 900 mV | 895.3 |  | 904.7 | mV | [5 empty] | DCV | 4 | DC | 0.001 | 0.9 | 0.9 | 20M | 50
| Ch 4, 50 Ohm, 20 MHz BW | 1 mV/div | 0 mV | -0.2 |  | 0.2 | mV | [5 empty] | DCV | 4 | DC | 0.001 | 0 | 0 | 20M | 50
| Ch 4, 50 Ohm, 20 MHz BW | 1 mV/div | -900 mV | -904.7 |  | -895.3 | mV | [5 empty] | DCV | 4 | DC | 0.001 | -0.9 | -0.9 | 20M | 50
| Ch 4, 50 Ohm, 20 MHz BW | 100 mV/div | 3 V | 2.97 |  | 3.03 | V | [5 empty] | DCV | 4 | DC | 0.1 | 3 | 3 | 20M | 50
| Ch 4, 50 Ohm, 20 MHz BW | 100 mV/div | 0 V | -0.02 |  | 0.02 | V | [5 empty] | DCV | 4 | DC | 0.1 | 0 | 0 | 20M | 50
| Ch 4, 50 Ohm, 20 MHz BW | 100 mV/div | -3 V | -3.03 |  | -2.97 | V | [5 empty] | DCV | 4 | DC | 0.1 | -3 | -3 | 20M | 50
6 | Ch 5, 1M Ohm, 20 MHz BW | 1 mV/div | 900 mV | 895.3 |  | 904.7 | mV | [5 empty] | DCV | 5 | DC | 0.001 | 0.9 | 0.9 | 20M | 1M
6 | Ch 5, 1M Ohm, 20 MHz BW | 1 mV/div | 0 mV | -0.2 |  | 0.2 | mV | [5 empty] | DCV | 5 | DC | 0.001 | 0 | 0 | 20M | 1M
6 | Ch 5, 1M Ohm, 20 MHz BW | 1 mV/div | -900 mV | -904.7 |  | -895.3 | mV | [5 empty] | DCV | 5 | DC | 0.001 | -0.9 | -0.9 | 20M | 1M
6 | Ch 5, 1M Ohm, 20 MHz BW | 100 mV/div | 1.0 V | 0.975 |  | 1.025 | V | [5 empty] | DCV | 5 | DC | 0.1 | 1 | 1 | 20M | 1M
6 | Ch 5, 1M Ohm, 20 MHz BW | 100 mV/div | 0 V | -0.02 |  | 0.02 | V | [5 empty] | DCV | 5 | DC | 0.1 | 0 | 0 | 20M | 1M
6 | Ch 5, 1M Ohm, 20 MHz BW | 100 mV/div | -1.0 V | -1.025 |  | -0.975 | V | [5 empty] | DCV | 5 | DC | 0.1 | -1 | -1 | 20M | 1M
6 | Ch 5, 1M Ohm, 20 MHz BW | 500 mV/div | 9.0 V | 8.855 |  | 9.145 | V | [5 empty] | DCV | 5 | DC | 0.5 | 9 | 9 | 20M | 1M
6 | Ch 5, 1M Ohm, 20 MHz BW | 500 mV/div | 0 V | -0.1 |  | 0.1 | V | [5 empty] | DCV | 5 | DC | 0.5 | 0 | 0 | 20M | 1M
6 | Ch 5, 1M Ohm, 20 MHz BW | 500 mV/div | -9.0 V | -9.145 |  | -8.855 | V | [5 empty] | DCV | 5 | DC | 0.5 | -9 | -9 | 20M | 1M
6 | Ch 5, 1M Ohm, 20 MHz BW | 5 V/div | 10 | 8.95 |  | 11.05 | V | [5 empty] | DCV | 5 | DC | 5 | 10 | 10 | 20M | 1M
6 | Ch 5, 1M Ohm, 20 MHz BW | 5 V/div | 0 V | -1 |  | 1 | V | [5 empty] | DCV | 5 | DC | 5 | 0 | 0 | 20M | 1M
6 | Ch 5, 1M Ohm, 20 MHz BW | 5 V/div | -10 V | -11.05 |  | -8.95 | V | [5 empty] | DCV | 5 | DC | 5 | -10 | -10 | 20M | 1M
6 | Ch 5, 50 Ohm, 20 MHz BW | 1 mV/div | 900 mV | 895.3 |  | 904.7 | mV | [5 empty] | DCV | 5 | DC | 0.001 | 0.9 | 0.9 | 20M | 50
6 | Ch 5, 50 Ohm, 20 MHz BW | 1 mV/div | 0 mV | -0.2 |  | 0.2 | mV | [5 empty] | DCV | 5 | DC | 0.001 | 0 | 0 | 20M | 50
6 | Ch 5, 50 Ohm, 20 MHz BW | 1 mV/div | -900 mV | -904.7 |  | -895.3 | mV | [5 empty] | DCV | 5 | DC | 0.001 | -0.9 | -0.9 | 20M | 50
6 | Ch 5, 50 Ohm, 20 MHz BW | 100 mV/div | 3 V | 2.97 |  | 3.03 | V | [5 empty] | DCV | 5 | DC | 0.1 | 3 | 3 | 20M | 50
6 | Ch 5, 50 Ohm, 20 MHz BW | 100 mV/div | 0 V | -0.02 |  | 0.02 | V | [5 empty] | DCV | 5 | DC | 0.1 | 0 | 0 | 20M | 50
6 | Ch 5, 50 Ohm, 20 MHz BW | 100 mV/div | -3 V | -3.03 |  | -2.97 | V | [5 empty] | DCV | 5 | DC | 0.1 | -3 | -3 | 20M | 50
6 | Ch 6, 1M Ohm, 20 MHz BW | 1 mV/div | 900 mV | 895.3 |  | 904.7 | mV | [5 empty] | DCV | 6 | DC | 0.001 | 0.9 | 0.9 | 20M | 1M
6 | Ch 6, 1M Ohm, 20 MHz BW | 1 mV/div | 0 mV | -0.2 |  | 0.2 | mV | [5 empty] | DCV | 6 | DC | 0.001 | 0 | 0 | 20M | 1M
6 | Ch 6, 1M Ohm, 20 MHz BW | 1 mV/div | -900 mV | -904.7 |  | -895.3 | mV | [5 empty] | DCV | 6 | DC | 0.001 | -0.9 | -0.9 | 20M | 1M
6 | Ch 6, 1M Ohm, 20 MHz BW | 100 mV/div | 1.0 V | 0.975 |  | 1.025 | V | [5 empty] | DCV | 6 | DC | 0.1 | 1 | 1 | 20M | 1M
6 | Ch 6, 1M Ohm, 20 MHz BW | 100 mV/div | 0 V | -0.02 |  | 0.02 | V | [5 empty] | DCV | 6 | DC | 0.1 | 0 | 0 | 20M | 1M
6 | Ch 6, 1M Ohm, 20 MHz BW | 100 mV/div | -1.0 V | -1.025 |  | -0.975 | V | [5 empty] | DCV | 6 | DC | 0.1 | -1 | -1 | 20M | 1M
6 | Ch 6, 1M Ohm, 20 MHz BW | 500 mV/div | 9.0 V | 8.855 |  | 9.145 | V | [5 empty] | DCV | 6 | DC | 0.5 | 9 | 9 | 20M | 1M
6 | Ch 6, 1M Ohm, 20 MHz BW | 500 mV/div | 0 V | -0.1 |  | 0.1 | V | [5 empty] | DCV | 6 | DC | 0.5 | 0 | 0 | 20M | 1M
6 | Ch 6, 1M Ohm, 20 MHz BW | 500 mV/div | -9.0 V | -9.145 |  | -8.855 | V | [5 empty] | DCV | 6 | DC | 0.5 | -9 | -9 | 20M | 1M
6 | Ch 6, 1M Ohm, 20 MHz BW | 5 V/div | 10 | 8.95 |  | 11.05 | V | [5 empty] | DCV | 6 | DC | 5 | 10 | 10 | 20M | 1M
6 | Ch 6, 1M Ohm, 20 MHz BW | 5 V/div | 0 V | -1 |  | 1 | V | [5 empty] | DCV | 6 | DC | 5 | 0 | 0 | 20M | 1M
6 | Ch 6, 1M Ohm, 20 MHz BW | 5 V/div | -10 V | -11.05 |  | -8.95 | V | [5 empty] | DCV | 6 | DC | 5 | -10 | -10 | 20M | 1M
6 | Ch 6, 50 Ohm, 20 MHz BW | 1 mV/div | 900 mV | 895.3 |  | 904.7 | mV | [5 empty] | DCV | 6 | DC | 0.001 | 0.9 | 0.9 | 20M | 50
6 | Ch 6, 50 Ohm, 20 MHz BW | 1 mV/div | 0 mV | -0.2 |  | 0.2 | mV | [5 empty] | DCV | 6 | DC | 0.001 | 0 | 0 | 20M | 50
6 | Ch 6, 50 Ohm, 20 MHz BW | 1 mV/div | -900 mV | -904.7 |  | -895.3 | mV | [5 empty] | DCV | 6 | DC | 0.001 | -0.9 | -0.9 | 20M | 50
6 | Ch 6, 50 Ohm, 20 MHz BW | 100 mV/div | 3 V | 2.97 |  | 3.03 | V | [5 empty] | DCV | 6 | DC | 0.1 | 3 | 3 | 20M | 50
6 | Ch 6, 50 Ohm, 20 MHz BW | 100 mV/div | 0 V | -0.02 |  | 0.02 | V | [5 empty] | DCV | 6 | DC | 0.1 | 0 | 0 | 20M | 50
6 | Ch 6, 50 Ohm, 20 MHz BW | 100 mV/div | -3 V | -3.03 |  | -2.97 | V | [5 empty] | DCV | 6 | DC | 0.1 | -3 | -3 | 20M | 50
Analog Bandwidth
| Test Conditions | Rise Time (measured) | Lower Limit | Bandwidth (Calculated) | Units | Pass/Fail | [8 empty] | Timebase (ns)
| Channel 1 |  | 1000 |  | MHz | [7 empty] | RISE | 1 | 1 | 50
| Channel 2 |  | 1000 |  | MHz | [7 empty] | RISE | 2 | 1 | 50
| Channel 3 |  | 1000 |  | MHz | [7 empty] | RISE | 3 | 1 | 50
| Channel 4 |  | 1000 |  | MHz | [7 empty] | RISE | 4 | 1 | 50
6 | Channel 5 |  | 1000 |  | MHz | [7 empty] | RISE | 5 | 1 | 50
6 | Channel 6 |  | 1000 |  | MHz | [7 empty] | RISE | 6 | 1 | 50
Sample Rate and Delay Time Accuracy
| Test Conditions | Lower Limit | Measured | Upper Limit | Units | Pass/Fail
| Reference Oscillator | 10 |  | 10 | MHz
Auxiliary (Trigger) Output High Level
| Test Conditions | Impedance | Lower Limit | Measured | Units | Pass/Fail
| Trigger Output High | 1M Ohm | 2.5 |  | V
| Trigger Output High | 50 Ohm | 1 |  | V
Auxiliary (Trigger) Output Low Level
| Test Conditions | Impedance | Measured | Upper Limit | Units | Pass/Fail
| Trigger Output Low | 1M Ohm |  | 0.7 | V
| Trigger Output Low | 50 Ohm |  | 0.25 | V
End of Report